## Annex II. Tech.spec.
Annex II. OF TENDER DOCUMENTATION: Detailed Technical Specifications |  |  | Procurement subject description : AUTONOMOUS PASSIVE ACOUSTIC RECORDER (APAR), A PISTONPHONE AND ONSITE TRAINING
Blue World Institute of Marine Research and Conservation
NOTE FOR THE SUPPLIERS (TENDERERS): • The Tenderer must complete all the requested columns in this Annex; • The Tenderer must fill the name / model and manufacturer of the offered item; • In the column OFFERED SPECIFICATIONS the Tenderer must point out the exact specifications of the supply goods, that is, state the specific values of the product it offers • Where required, the Tenderer is obliged to state the number/quantity/value of the offered product without specifying the minimum, maximum, up tp, from and similar • All requested specifications are mandatory and tenderers are required to meet minimal technical specifications. • If the Tenderer leaves the column name/model and manufacturer blank the Contracting Authority may request further clarification from the Tenderer
No. | Ref | REQUESTED SPECIFICATION | OFFERED SPECIFICATIONS
1. | 1. AUTONOMOUS PASSIVE ACOUSTIC RECORDER
Name/model and manufacturer
| A | Recording file format: Waveform Audio File Format (.wav)
| B | Sensitivity: In the range -165 to -185 dB re 1 V/μPa
| C | Dynamic range: Minimum 16 bit |  | Note: Weighted points for added technical requirements(listed in the table below)
| D | Frequency response: 10 Hz – 10 kHz flat within ±2 dB, 10 Hz – 20kHz flat within ±3 dB
| E | Directionality: Omnidirectional to within +/- 1 dB up to 20 kHz horizontal, and to within +/- 2 dB in vertical
| F | Sampling rate: Minimum 44 kHz in the overall dynamic range |  | Note: Weighted points for added technical requirements(listed in the table below)
| G | Data storage: Minimum 512 GB (SD cards or SSD only) |  | Note: Weighted points for added technical requirements(listed in the table below)
| H | System self-noise: Better than 58 dB re 1 μPa2/Hz at 63 Hz; Better than 53 dB re 1 μPa2/Hz at 125 Hz
| J | Battery life: Alkaline: not less than 25 days Lithium: not less than 75 days Continuous recording 16 bit resolution, 24 kHz sampling |  | Note: Weighted points for added technical requirements(listed in the table below)
| K | Calibration: The system has to be fully calibrated (hydrophone sensitivity and directivity, gain and self noise) and the calibration documented for each instrument (no generic data)
| L | Programmable or switchable Input gain (has to enable various sensitivities within range required)
| M | Power up/down without opening of housing: Yes
| N | Start/stop recording without opening of housing: Yes
| P | Maximum operating depth: 500 m
| Q | Temperature range: -5°C to +40°C
| R | Communication interface for control and setting up of system parameters: Yes
| S | Mooring frame: Frame to enable attachment of deployment equipment (flotation, acoustic releasers, moorings etc.) to the instrument
| T | Hydrophone protection: Mechanical protection of hydrophone against physical damage (e.g. cage)
| U | Possibility of start stop recording (duty cycling, on/off scheduling): Yes
| V | Dimensions (with mooring frame and batteries): Diameter no more than 220 mm Length no more than 1100 mm Weight in water no more than 20 kg
| W | Possibility to use both alkaline and lithium batteries: Yes
| AB | Warranty Minimum 1 year
2. | 2. PISTONPHONE
Name of the item
| AC | with coupler for the hydrophone quoted
Value added technical requirements
| Ref | Additional technical requirement | Weighting points
| C1 | Dynamic range 16 and 24 bit | Maximum added points: 3 points If 16 bit= 0 points (16 bit is mandatory according to the Mandatory or minimum technical requirements, point C - Dynamic range) If 24 bit or more = 3 points
| F1 | Additional sampling rate | Maximum added points: 3 points If 44 kHz= 0 points (44 kHz is mandatory according to the Mandatory or minimum technical requirements, point F – sampling rate) If additional sampling rate to the one of 44 kHz is offered = 3 points
| G1 | Additional data storage (more than 512 GB) | Maximum added points: 3 points If 512 GB= 0 points (512 GB is mandatory according to the Mandatory or minimum technical requirements, point G-data storage) If 1 TB or more = 3 points
| J1 | Increased battery life (alkaline) | Maximum added points: 3 points If equal to 25 days = 0 points (25 days is mandatory according to the Mandatory or minimum technical requirements, point J-battery life, alkaline) If equal to 30 days or more = 3 points
| J2 | Increased battery life (lithium) | Maximum added points: 3 points If equal to 75 days = 0 points (75 days is mandatory according to the Mandatory or minimum technical requirements, point J-battery life, lithium) If equal to 90 or more = 3 points
Note on A32: [Komentar u obliku niti] Vaša verzija programa Excel omogućuje vam da pročitate ovaj komentar u obliku niti, no sve će izmjene biti uklonjene ako datoteka bude otvorena u novijoj verziji programa Excel. Saznajte više: https://go.microsoft.com/fwlink/?linkid=870924 Komentar: Obavezno dodajte i točku 3. Training.
Note on C39: [Komentar u obliku niti] Vaša verzija programa Excel omogućuje vam da pročitate ovaj komentar u obliku niti, no sve će izmjene biti uklonjene ako datoteka bude otvorena u novijoj verziji programa Excel. Saznajte više: https://go.microsoft.com/fwlink/?linkid=870924 Komentar: Ovdje je savjet korigrati tablicu prema sugestiji u Dokumentaciji.