PRILOG II

TEHNIČKE SPECIFIKACIJE /

ANNEX II

TECHNICAL SPECIFICATIONS

Tehničke specifikacije sastavni su dio ovog Poziva na dostavu ponuda te se iste nalaze u privitku./ Technical specifications are integral part of this Tender documentation and can be found as an attachments to this Tender documentation.

Ponuda mora zadovoljiti sve tražene tehničke karakteristike, te obuhvatiti sve stavke iz tehničke specifikacije koje predstavljaju minimalne tehničke karakteristike (nije dozvoljena modifikacija istih od strane Ponuditelja). Ponuditelji mogu nuditi i proizvode koji ispunjavaju i više standarde od postavljenih minimalnih./ Tenderer must satisfy all the required technical specifications and include all items in the technical specification that represent the minimum technical specifications. Tenderers may also offer products that meet higher standards, i.e. technical specifications, than the set minimum specifications (the Tenderer is not allowed to modify technical specifications).

Ponuditelj OBAVEZNO POPUNJAVA stupac 3. Ponuditelj je dužan ispuniti svaku stavku u stupcu na način da „Ponuđene specifikacije“ ispuni traženim podatkom, odnosno vrijednošću koju nudi, odnosno sa „DA“ ili „NE“ potvrdi da predmet nabave zadovoljava traženu tehničku karakteristiku. Kako bi se ponuda smatrala valjanom, ponuđeni predmet nabave mora zadovoljiti sve što je traženo u obrascu Tehničkih specifikacija./ Tenderer MUST COMPLETE Column 3. Tenderer is obliged to fill in each item in the column by filling in the "Offered Specifications" with the required data, ie the value it offers, or with "YES" or "NO" confirms that the procurement subject meets the required technical specifications. In order for the Tenders to be considered valid, the economic operator must satisfy all required in the Technical Specifications form.

Za sve stavke navedene u tehničkim specifikacijama u kojima se možebitno traži ili navodi marka, norma te standardi, patent, tip ili određeno podrijetlo ponuditelj može ponuditi „jednakovrijedno“ svemu traženom ili navedenom./ For all items listed in the technical specifications in which the designation, norms and standards, patent, type or specific origin may be requested or indicated, the tenderer may offer an "equivalent" to all required or specified.

Ponuditelj ispunjava tehničke specifikacije samo za grupu predmeta nabave za koju podnosi ponudu./ The tenderer fills in the technical specifications only for the group of procurement for which it is submitting a bid.

Stupac « Bilješke, napomene, reference na tehničku dokumentaciju » Ponuditelj može popuniti ukoliko smatra potrebnim./The Tenderer  may fill in the column "Notes, remarks, references to technical documentation" if he deems it necessary.

GRUPA 1 Jedinice za čišćenje, elektropoliranje, pasiviranje

GROUP 1 Cleaning, electropolishing and passivation equipment

| - Ultrazvučni sustavi za prethodno čišćenje Ultrasonic pre-cleaning |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Nehrđajući čelik, debljine minimalno 2 mm, elektrolitički polirani Stainless Steel, thickness 2mm, electrolytic polished |  |  |
| Geometrija / Izvedba Geometry / Execution | - Sigurnosni preljev na stražnjoj strani spremnika, s priključkom na odvod za čišćenje Safety overflow on the back side of the tank, with connection to drain for the cleaning media - Nagnuto dno komore za jamčenje potpune drenaže spremnika Sloped chamber bottom to guarantee the complete drainage of the tank - Toplinska izolacija Thermal insulation |  |  |
| Opskrba vodom Water-supply | - Automatsko punjenje/dopunjavanje spremnika Automatic filling / refilling of the tank |  |  |
| Kontrola razine Level control | - Kontrola razine tekućine za zaštitu ultrazvučne opreme od grijanja u slučaju rada na suho Liquid level control for protection of the US equipment and heating against dry running |  |  |
| Odvodnja Draining | - Ručno pražnjenje spremnika s kuglastim ventilom Manual draining of the tank with ball valve |  |  |
| Ultrazvučni oprema Generator Ultrasonic equipment Generator | - 2 ultrazvučna generatora kontrolirana mikroprocesorom - svaki s konstantnom izlaznom snagom minimalno 600W i vršnom izlaznom snagom minimalno 1200W 2 Micro-processor controlled ultrasonic generators - each with constant output power of minimum 600W, peak output power minimum 1200W - Frekvencija 27/80 kHz, automatsko pomicanje frekvencije Frequency 27/80kHz, automatic frequency-sweeping |  |  |
| Ultrazvučni oprema Pretvornik Ultrasonic equipment Transducer | - Ultrazvučni pretvarači izravno pričvršćeni na vanjsku stranu spremnika - na dnu i na jednoj strani spremnika, svaki s konstantnom izlaznom snagom od minimalno 600 W, vršnom izlaznom snagom minimalno 1'200 W Ultrasonic transducers directly bonded onto the outside of the tank - on the bottom and on one side of the tank, each with a constant output of minimum 600W, peak output power minimum 1200W - Frekvencija 27/80 kHz Frequency 27/80kHz |  |  |
| Grijanje Heating | - Ukupna snaga od minimalno 2kW Total power of minimum 2 kW - Vrijeme zagrijavanja za Δt od 30°C 75 minuta minimalno Heating time for a Δt of 30°C 75 minutes minimum |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-60°C Electronic temperature control and regulation 20-60° C |  |  |
| Poklopac Lid | - Ručni poklopac, bez šarki Manual non-hinged lid |  |  |

| 1.2. Ultrazvučno ispiranje Ultrasonic Rinsing |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Nehrđajući čelik, debljine minimalno 2 mm, elektrolitički polirani Stainless Steel, minimal thickness 2mm, electrolytic polished |  |  |
| Geometrija / Izvedba Geometry / Execution | - Spremnik za ispiranje, s preljevom vode spojenim na odvod vode Rinsing tank, with water overflow connected to drain for rinsing water - Nagnuto dno komore kako bi se osigurala potpuna drenaža radne komore Inclined chamber bottom to ensure a complete drainage of the working chamber - Sigurnosni preljev spojen na odvod za ispiranje Safety overflow connected to drain for rinsing media |  |  |
| Opskrba vodom Water-supply | - Automatsko punjenje/dopunjavanje spremnika Automatic filling / refilling of the tank |  |  |
| Kontrola razine Level control | - Kontrola razine tekućine za zaštitu ultrazvučne opreme od grijanja u slučaju rada na suho Liquid level control for protection of the US equipment and heating against dry running. |  |  |
| Odvodnja Draining | - Ručno pražnjenje spremnika s kuglastim ventilom Manual draining of the tank with ball valve |  |  |
| Ultrazvučni oprema Generator Ultrasonic equipment Generator | - 2 ultrazvučna generatora kontrolirana mikroprocesorom - svaki s konstantnom izlaznom snagom od minimalno 600W, vršnom izlaznom snagom minimalno 1200W 2 Micro-processor controlled ultrasonic generators - each with constant output power of minimum 600W, peak output power minimum 1'200W - Frekvencija 27/80 kHz, automatsko pomicanje frekvencije Frequency 27/80kHz, automatic frequency-sweeping |  |  |
| Ultrazvučni oprema Pretvornik Ultrasonic equipment Transducer | - Ultrazvučni pretvarači izravno pričvršćeni na vanjsku stranu spremnika - na dnu i na jednoj strani spremnika, svaki s konstantnom izlaznom snagom od minimum 600 W, vršnom izlaznom snagom minimum 1200 W Ultrasonic transducers directly bonded onto the outside of the tank - on the bottom and on one side of the tank, each with a constant output power of minimum 600W, peak output power minimum 1'200W - Frekvencija 27/80 kHz Frequency 27/80kHz |  |  |
| Grijanje Heating | - Ukupna snaga od minimalno 2kW Total power of 2 kW minimum - Vrijeme zagrijavanja za Δt od 30°C 75 minuta minimum Heating time for a Δt of 30°C 75 minutes minimum |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-60°C Electronic temperature control and regulation 20-60° C |  |  |
| Poklopac Lid | - Ručni poklopac, bez šarki Manual non-hinged lid |  |  |

| 1.3. Spremnik 3: Elektropoliranje Tank 3: Electrolytic polishing |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Komora od PP, debljine 10 mm minimalno Chamber made of PP, thickness 10 mm minimum |  |  |
| Geometrija / Izvedba Geometry / Execution | - Spremnik s obodnim pojačanjem, donji nosač postavljen u okvir od nehrđajućeg čelika Tank with circumferential reinforcement, bottom support placed in stainless-steel frame - Integrirano horizontalno kretanje - Horizontalno oscilirajuće kretanje robe - Max. Nosivost 5 kg Integrated Horizontal movement - Horizontal oscillating goods movement - Max. Payload 5 kg - Dno komore s držačem za pumpu kako bi se osiguralo potpuno čišćenje Chamber bottom with pump-dwell to ensure a complete cleanout |  |  |
| Touch Panel | - Upravljanje ispravljačem pomoću PLC-a i preko touch panela Control of the rectifier with PLC and over a touch panel - Kontrola procesa konstantne struje Constant current process control - Zvučni alarm na kraju definiranog procesa Acoustic alarm at the finish of the defined process - Upravljački elementi na prednjoj strani spremnika za uključivanje i isključivanje ispravljača Control elements on the front of the tank to switch on and off of the rectifier |  |  |
| Katoda Cathode | - Visokostrujna kontaktna sjedala na vrhu spremnika za anodni kontakt i držanje nosača proizvoda High current contact saddles on top of the tank for the anodic contacting and holding of the product carriers - Spajanje na ispravljač fleksibilnim, izoliranim bakrenim kabelom Connection to the rectifier by flexible, insulated copper cable - Katodni listovi od nehrđajućeg čelika na dvije duge strane spremnika Stainless Steel Cathode sheets on two long sides of the tank |  |  |
| Ispravljač Rectifier | - 15 V/600 A minimalno 15 V/600 A minimum - Stabilizacija napona i struje Stabilization of voltage and amperage - Aktivno hlađenje zrakom s integriranim ventilatorom Active air cooling with integrated fan - Povremeno i nasumično odstupanje < 3% u ukupnom radnom rasponu Periodic and random deviation < 3% over the total operating range - Točnost upravljanja u stacionarnom stanju < 1% Steady-state control accuracy < 1% - Integriran u središnji upravljački ormar Integrated in central control cabinet |  |  |
| Punjenje Filling | - Ručno punjenje materijala Manual filling of the media |  |  |
| Kontrola razine Level Control | - Integrirane kontrole razine tekućine za zaštitu grijanja i kontrolu funkcije ponovnog punjenja Integrated liquid level controls to protect the heating and control the refilling function |  |  |
| Cirkulacija materijala Media Circulation | - Krug integrirane cirkulacijske crpke od PVDF-a (Poliviniliden fluorid) Integrated Circulation Pump circuit made of PVDF |  |  |
| Grijanje Heating | - Potopni grijaći elementi od porculana ukupne snage 2 kW minimalno Submersible porcelain-heating elements with total power of 2 kW minimum - Integrirana zaštita od pregrijavanja Integrated protection against overheating - Elektronska kontrola i regulacija temperature (30 - 80°C) Electronic temperature control and regulation (30 - 80°C) |  |  |
| Hlađenje Cooling | - Integrirani sustav hlađenja s radijatorom u komori Integrated cooling system with radiator in the chamber - Priključak na krug rashladne vode na licu mjesta Connection with the onsite cooling water loop |  |  |
| Regulacija temperature Temperature Regulation | - Regulacija I kontrola temperature s termičkim isključivanjem Temperature-control and regulation with thermal cut-out - 2 neovisna mjerača temperature unutar radne komore 2 independent temperature gauges within the working chamber - Tjedni timer/mjerač vremena i kontrola temperature integrirana u PLC (industrijsko računalo) Weekly timer and control of the temperature integrated in the PLC |  |  |
| Ispuštanje Exhausting | - Sustav za odsis pare s dvije strane ispod poklopca s priključkom na središnji odsisni sustav na stražnjoj strani jedinice s integriranim ventilom Vapor exhausting system on two sides underneath the lid with connection to the central exhausting system on the backside of the unit with integrated valve |  |  |
| Automatski klizni poklopac Automatic sliding lid | - Automatsko klizno otvaranje poklopca prema stražnjoj strani, pneumatski pogon Automatic sliding lid opening to the rear, pneumatically driven |  |  |

| 1.4. Ispiranje nakon potapanja – Voda obogaćena elektrolitom Post Dip Rinsing – Water enriched with electrolyte |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Komora od PP, debljine 10 mm minimalno Chamber made of PP, thickness 10 mm minimum |  |  |
| Geometrija / Izvedba Geometry / Execution | - Ojačani spremnik postavljen u okvir od nehrđajućeg čelika Reinforced tank placed in stainless-steel frame - Nagnuto dno komore za jamčenje potpune drenaže spremnika Sloped chamber bottom to guarantee the complete drainage of the tank |  |  |
| Punjenje / Dopunjavanje Filling / Refilling | - Ručno punjenje Manual filling |  |  |
| Ispuštanje Exhausting | - Sustav za odsis pare s dvije strane ispod poklopca s priključkom na središnji odsisni sustav na stražnjoj strani jedinice s integriranim ventilom Vapor exhausting system on two sides underneath the lid with connection to the central exhausting system on the backside of the unit with integrated valve |  |  |
| Grijanje Heating | - Potopni PTFE grijaći elementi ukupne snage minimalno 2 kW Submersible PTFE heating elements with total power of 2 kW minimum - Integrated protection against overheating Integrirana zaštita od pregrijavanja |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-60° Electronic temperature control and regulation 20-60° - Tjedni mjerač/timer vremena za programiranje vremena grijanja Weekly timer for programming of the heating times |  |  |
| Odvor Drain | - Ručno ispuštanje spremnika s kuglastim ventilom Manual drain of the tank with ball valve - Spojen na odvod za čišćenje Connected to drain for Cleaning Media |  |  |
| Kontrola razine Level Control | - Integrirana kontrola razine tekućine Liquid level control integrated - Kontrola razine uz pomoć uzlaznih vodova smještenih s na vanjskoj strani spremnika Level control with riser-tube on the outside of the tank |  |  |
| Poklopac Lid | - Ručni poklopac bez šarki ili klizni poklopac od PP Manual non-hinged or sliding lid in PP |  |  |

| 1.5. Spremnik 5: Protok vode kroz reverznu osomozu – Ispiranje “CR2” Tank 5: Tap-/RO-Water Flow-Rinsing “CR2” |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Komora od PP, debljine 10 mm minimalno Chamber made of PP, thickness 10 mm minimum |  |  |
| Geometrija / Izvedba Geometry / Execution | - Spremnik za ispiranje, s preljevom vode spojenim na odvod za ispiranje materijala Rinsing tank, with water overflow connected to drain for rinsing media - Nagnuto dno komore kako bi se osigurala potpuna drenaža radne komore Inclined chamber bottom to ensure a complete drainage of the working chamber |  |  |
| Protočno ispiranje Flow rinsing | - Ručni gumb na prednjoj strani spremnika za pokretanje i zaustavljanje protoka vode Manual push button on the front of the tank to start and stop the water-flow |  |  |
| Opskrba vodom Water-supply | - Automatsko punjenje/dopunjavanje spremnika, elektromagnetni ventil spojen na slavinu/reverznu osmozu Automatic filling/ refilling of the tank, solenoid valve connected to Tap-/RO-Water - Statičko ispiranje ili ispiranje alternativnog protoka unosom vode s ručnom regulacijom protoka Static rinsing or alternative flow rinsing by water intake with manual flow regulation |  |  |
| Kontrola razine Level Control | - Kontrola razine tekućine za zaštitu od grijanja u slučaju rada na suho Liquid level control for protection of the heating against dry running |  |  |
| Odvod Drain | - Ručno pražnjenje spremnika s kuglastim ventilom Manual draining of the tank with ball valve - Spojen na središnji odvod za ispiranje materijala Connected to the central drain for the Rinsing Media |  |  |
| Grijanje Heating | - Potopni PTFE grijaći elementi ukupne snage 2 kW minimalno Submersible PTFE heating elements with total power of 2 kW minimum - Integrirana zaštita od pregrijavanja Integrated protection against overheating |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-60° Electronic temperature control and regulation 20-60° - Tjedni mjerač/timer vremena za programiranje vremena grijanja Weekly timer for programming of the heating times |  |  |
| Poklopac Lid | - Ručni poklopac bez šarki ili klizni poklopac od PP Manual non-hinged or sliding lid in PP |  |  |

| 1.6. Spremnik 6: Pasiviranje “Dušična kiselina 2” Tank 6: Passivation “Nitric 2” |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - PVDF spremnik, debljina minimalno 10 mm PVDF tank, thickness minimum 10mm |  |  |
| Geometrija / Izvedba Geometry / Execution | - Nagnuto dno komore za jamčenje potpune drenaže spremnika Sloped chamber bottom to guarantee the complete drainage of the tank |  |  |
| Cirkulacija medija Media Circulation | - Integrirana cirkulacija za kontinuirano kretanje i raspodjelu temperature medija za pasiviranje Integrated circulation for continuous movement and temperature distribution of the passivation media - Pumpa otporna na kemikalije s protokom od 10-40l/min Chemical resistant pump with a flow rate of 10-40l/min - Ručna kontrola protoka s kuglastim ventilom, nadzor mjeračem protoka s plovkom Manual flow control with ball valve, monitoring with float flow meter - Cijevi otporne na kemikalije od PVDF-a Chemical resistant piping in PVDF |  |  |
| Opskrba vodom Water-supply | - Ručno punjenje Manual filling |  |  |
| Doziranje sredstva za pasiviranje Dosing passivation agent | - Ručno doziranje sredstva za pasiviranje Manual dosing of the passivation agent |  |  |
| Primarni ispušni sustav Primary exhaust system | - Usisni otvori ispod automatskog poklopca Suction openings below the automatic lid |  |  |
| Sekundarni ispušni sustav Secondary exhaust system | - Usisni okvir iznad automatskog poklopca Suction frame above the automatic lid - Otvori 3-4 strane spremnika, konstantno usisavanje oko ovratnika spremnika Openings 3-4 sides of the tank, constant suction all around the tank collar |  |  |
| Kontrola razine Level Control | - Kontrola razine tekućine za zaštitu od grijanja u slučaju rada na suho Liquid level control for protection of the heating against dry running |  |  |
| Odvod Drain | - Ručno pražnjenje spremnika s bubanj pompom Manual draining of the tank with drum pump |  |  |
| Grijanje Heating | - Potopni PTFE grijaći elementi ukupne snage 2 kW minimalno Submersible PTFE heating elements with total power of 2 kW minimum - Integrirana zaštita od pregrijavanja Integrated protection against overheating |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-60° Electronic temperature control and regulation 20-60° - Tjedni mjerač/timer vremena za programiranje vremena grijanja Weekly timer for programming of the heating times |  |  |
| Poklopac Lid | - Automatski klizni poklopac, pneumatski pogon Automatic sliding lid, pneumatically driven |  |  |

| 1.7. Spremnik 7: Ispiranje nakon pasiviranja Tank 7: Post Dip Passivation Rinsing |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Komora od PP, debljine 10 mm minimalno Chamber made of PP, Thickness 10 mm minimum |  |  |
| Geometrija / Izvedba Geometry / Execution | - Ojačani spremnik postavljen u okvir od nehrđajućeg čelika Reinforced tank placed in stainless-steel frame - Nagnuto dno komore za jamčenje potpune drenaže spremnika Sloped chamber bottom to guarantee the complete drainage of the tank |  |  |
| Punjenje / Dopunjavanje Filling / Refilling | - Ručno punjenje Manual filling |  |  |
| Ispuštanje Exhausting | - Sustav za odsis pare s dvije strane ispod poklopca s priključkom na središnji odsisni sustav na stražnjoj strani jedinice s integriranim ventilom Vapor exhausting system on two sides underneath the lid with connection to the central exhausting system on the backside of the unit with integrated valve |  |  |
| Grijanje Heating | - Potopni PTFE grijaći elementi ukupne snage 2 kW minimalno Submersible PTFE heating elements with total power of 2 kW minimum - Integrirana zaštita od pregrijavanja Integrated protection against overheating |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-60° Electronic temperature control and regulation 20-60° - Tjedni mjerač/timer vremena za programiranje vremena grijanja Weekly timer for programming of the heating times |  |  |
| Odvod Drain | - Ručni odvod spremnika s kuglastim ventilom, spojen na odvod za čišćenje materijala Manual drain of the tank with ball valve, connected to drain for Cleaning Media |  |  |
| Kontrola razine Level Control | - Integrirana kontrola razine tekućine Liquid level control integrated - Kontrola razine uz pomoć uzlaznih vodova smještenih na vanjskoj strani spremnika Level control with riser-tube on the outside of the tank |  |  |
| Poklopac Lid | - Ručni poklopac bez šarki ili klizni poklopac od PP Manual non-hinged or sliding lid in PP |  |  |

| 1.8. Spremnik 8+9: Deoinizirano protočno ispiranje Tank 8+9: DI Flow-Rinsing |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Nehrđajući čelik minimalne debljine 2 mm, elektrolitički polirani Stainless Steel, minimum thickness 2mm, electrolytic polished |  |  |
| Geometrija / Izvedba Geometry / Execution | - Kaskada za dvostruko ispiranje s pregradnim zidovima između dva položaja Double rinsing Cascade with partition walls between the two positions - Nagnuto dno komore za jamčenje potpune drenaže spremnika Sloped chamber bottom to guarantee the complete drainage of the tanks - Preljev u stražnjem dijelu spremnika, spojen na odvod za ispiranje materijala ili na DI-sustav za obradu vode Overflow in the back of the tank, connected to drain for rinsing media or to DI- Water treatment system |  |  |
| Protočno ispiranje Flow rinsing | - Ručni gumb na prednjoj strani spremnika za pokretanje i zaustavljanje protoka vode Manual push button on the front of the tank to start and stop the water-flow |  |  |
| Odvod Drain | - Ručno pražnjenje spremnika s kuglastim ventilom na odvod za ispiranje materijala Manual emptying of the tank with ball valve to drain for rinsing media |  |  |
| Grijanje Heating | - Mat grijači pričvršćeni na vanjsku stranu spremnika, ukupne snage 2 kW minimalno Mat heaters bonded onto the outside of the tank, total power of 2 kW minimum - Vrijeme zagrijavanja za Δt od 30°C 75 minuta minimalno Heating time for a Δt of 30°C 75 minutes minimum |  |  |
| Regulacija I kontrola temperature Temperature-control and regulation | - Elektronska kontrola i regulacija temperature 20-50° Electronic temperature control and regulation 20-50° - Tjedni mjerač/timer vremena za programiranje vremena grijanja Weekly timer for programming of the heating times |  |  |
| Poklopac Lid | - Ručni poklopac bez šarki ili klizni poklopac od nehrđajućeg čelika za oba spremnika Manual non-hinged or sliding lid in stainless steel for both tanks |  |  |

| 1.9. Spremnik 10: Sušenje vrućim zrakom Tank 10: Hot-Air-Drying |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Materijali/površine Materials / surfaces | - Nehrđajući čelik, minimalna debljina 2 mm Stainless steel, minimum thickness 2 mm |  |  |
| Geometrija / Izvedba Geometry / Execution | - Usmjerivači zraka na 1 strani komore za sušenje Air deflectors on 1 side of the drying chamber - Ispuštanje zraka na dnu komore Exhausting of the air on the bottom of the chamber - Toplinska izolacija 20 mm minimalno Thermal insulation 20 mm minimum |  |  |
| Ventilacija / Grijanje Ventilation / Heating | - 1 ventilator s 0,75 kW minimalno 1 fan with 0.75kW minimum - Grijaći element od nehrđajućeg čelika u struji zraka, 7,5kW minimalno Stainless steel heating element in the air flow, 7.5kW minimum - Usisavanje svježeg zraka iz okolnog prostora uređaja Suction of fresh air from the unit surrounding area |  |  |
| Kontrola temperature Temperature-control | - Elektronska kontrola i regulacija temperature 30 - 120°C Electronic temperature control and regulation 30 - 120° C - 2 neovisne temperaturne sonde u struji vrućeg zraka na mjestu ispuštanja 2 independent temperature probes in the hot air-stream at the exhausting site |  |  |

| 1.10. Mokri čistač zraka / Neutralizator Wet Air-Scrubber / Neutralizer |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specifications | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Djelovanje Operation | - Neutralizacija kontaminiranog ispušnog zraka od elektropoliranja Neutralization of contaminated exhaust air from Electro-Polishing |  |  |
| Materijali/površine Materials / surfaces | - Polypropylene Polipropilen |  |  |
| Geometrija / Izvedba Geometry / Execution | - Izravno spojen na odvodni sustav modula za elektropoliranje / pasiviranje Directly connected to exhausting System of the Electro polishing / Passivation modules - Integrirana apsorpcija kapljica i aerosola Integrated drop and aerosol absorption - Integrirano cjedilo u razdjelniku tekućine Integrated strainer in the liquid distributor - Cirkulacijska pumpa u PP Circulating pump in PP - Sonda za mjerenje pH i za kontrolu pH raspona Probe for the measurement of the pH and for the control of the pH-range - Ulaz i izlaz plina Ø 125 mm Gas-inlet and outlet Ø 125 mm - Prozirni proizvodni / kontrolni priključak Ø 200 mm Transparent manufacturing / control port Ø 200 mm |  |  |
| Kontaminacija sirovog plina Contamination raw gas | - Pare iz EP/PA otopine Vapor from EP/PA solution |  |  |
| Tekućina za neutralizaciju Neutralization liquid | - Natrijev hidroksid NaOH 10% Sodium hydroxide NaOH 10% |  |  |
| Kontrola razine Level Control | - Automatska kontrola razine tekućine i punjenje preko elektromagnetnog ventila Automatic liquid level control and refilling over solenoid valve - Kontrola razine uz pomoć uzlaznih vodova smještenih s na vrhu spremnika Level control with transparent riser-tube on top of the tank |  |  |
| Doziranje tekućine za neutralizaciju Dosing of neutralization liquid | - pH-mjerenje pH-measurement - Automatsko doziranje/dodavanje tekućine za neutralizaciju (NaOH) Automatic dosing/adding of the neutralization media (NaOH) |  |  |
| Radna temperatura Operating temperature | - 20 – max. 40 ° C 20 – max. 40 ° C |  |  |

| 1.11. Sustav za tretiranje Deionizirane vode DI-Water treatment System |  |  |  |
| --- | --- | --- | --- |
| Svojstvo Feature | Tražene specifikacije Required specifications | Ponuđene specifikacije Offered specification | Bilješke, napomene, reference na tehničku dokumentaciju Notes, remarks, references to technical documentation |
| 1. | 2. | 3. | 4. |
| Sustav System | - Priključak na centralni vodovod revezrne osmoze Connection to the central Town/RO-Water supply |  |  |
| Spremnik Tank | - Spremnik za držanje izrađen od nehrđajućeg čelika, elektrolitički poliran Holding tank made in stainless steel, electrolytic polished - Volumen 100 litara minimalno, ručni poklopac Volume 100 Liter minimum, manual lid - Sigurnosni preljev spojen na odvod Safety overflow connected to drain - Toplinska izolacija 10 mm minimalno Thermal insulation 10 mm minimum - Kontrola razine tekućine za zaštitu od grijanja u slučaju rada na suho Liquid level control for protection of the heating against dry running - Kontrola razine s usponom cijevi na vanjskoj strani spremnika Level control with riser-tube on the outside of the tank - Automatsko punjenje/dopunjavanje spremnika, elektromagnetni ventil spojen na centralno napajanje deionizirane vode Automatic filling / refilling of the tank, Solenoid valve connected with the central DI-Water supply |  |  |
| Grijanje Heating | - Mat grijači spojeni na vanjsku stranu spremnika, ukupne snage 6 kW minimalno Mat heaters bonded onto outside of the tank, Total power of 6 kW minimum |  |  |
| Kontrola temperature Temperature control | - Elektronska kontrola i regulacija temperature 30 - 60°C Electronic temperature control and regulation 30 - 60° C - Sonda za temperaturu na vanjskoj strani spremnika Temperature probe on the outside of the tank - Tjedni mjerač vremena za programiranje vremena grijanja Weekly timer for programming of the heating time |  |  |
| Pumpa Pump | - Cirkulacijska pumpa, max. protok 5-20l/min Circulating pump, max. flow-rate of 5-20l/min - Cijevi od nehrđajućeg čelika Piping in stainless steel - Rad 120-180l/h za kontinuiranu cirkulaciju deonizirane vode Operation 120-180l/h for continuous circulation of the DI-Water |  |  |
| Filter Filter | - Filter s aktivnim ugljenom Activated Carbon Filter - Mješovita ionska izmjena, 2 patrone s miješanom smolom, spojene u seriju Mixed Bed Ion Exchange, 2 cartridges with mixed bed resin, connected in series |  |  |
| UV dezinfekcija UV disinfection | - UV dezinfekcija, Integracija UV reaktora s UV lampom visokih performansi 70W, izlaz max. 600 l/h – UV disinfection, Integration of a UV reactor with high-performance UV lamp 70W, output max. 600 l/h – - UV nadzor s prikazom statusa, digitalnim brojačem radnih sati, odvodnim i odzračnim ventilima UV monitoring with status display, digital operating hours counter, drain and vent valves - Cijevi od nehrđajućeg čelika Piping in stainless steel |  |  |
| Mjerenje vodljivosti Conductivity Measurement | - Integrirano mjerenje vodljivosti Integrated measurement of the conductivity - Raspon mjerenja 0,05 – 20 µS/cm Measuring range 0.05 – 20 µS/cm |  |  |

____________________, ______/______/2022.	 M.P./Stamp			ZA PONUDITELJA / ON BEHALF OF TENDERER:

(mjesto, datum/place, date)

___________________________________________________

(ime, prezime i potpis ovlaštene osobe/ name and signature of authorized person